## Slide 1: Introduction to Data Independent Acquisition
Jarrett Egertson, Ph.D.
Staff Scientist / Consultant
MacCoss Lab

## Slide 2: Part 1: How Does DIA Work?

## Slide 3: Comprehensive MS/MS Acquisition
20 20 m/z-wide windows = 400 m/z
m/z
500
900
Notes: Switch to 20, put in some MS1 scans
Usually 1,000 – 5,000 peptides identified by MS/MS
Usually ~25,000 – 50,000 MS features detected
2-20% of Features Annotated
BUT Usually only ~70% overlap in annotated features between runs

## Slide 4: Comprehensive MS/MS Acquisition
Scan 1
20 20 m/z-wide windows = 400 m/z
m/z
500
900
Notes: Switch to 20, put in some MS1 scans
Usually 1,000 – 5,000 peptides identified by MS/MS
Usually ~25,000 – 50,000 MS features detected
2-20% of Features Annotated
BUT Usually only ~70% overlap in annotated features between runs

## Slide 5: Comprehensive MS/MS Acquisition
Scan 1
20 20 m/z-wide windows = 400 m/z
m/z
500
900
Scan 2
Notes: Switch to 20, put in some MS1 scans
Usually 1,000 – 5,000 peptides identified by MS/MS
Usually ~25,000 – 50,000 MS features detected
2-20% of Features Annotated
BUT Usually only ~70% overlap in annotated features between runs

## Slide 6: Comprehensive MS/MS Acquisition
Scan 1
20 20 m/z-wide windows = 400 m/z
m/z
500
900
Scan 2
Scan 3
Scan 4
Scan 5
Scan 6
Scan 7
Scan 20
…
Scan 21
Notes: Switch to 20, put in some MS1 scans
Usually 1,000 – 5,000 peptides identified by MS/MS
Usually ~25,000 – 50,000 MS features detected
2-20% of Features Annotated
BUT Usually only ~70% overlap in annotated features between runs

## Slide 7: Comprehensive MS/MS Acquisition
20 20 m/z-wide windows = 400 m/z
m/z
500
900
Time
~2 seconds
Precursor Filter
780-800 m/z
Orbitrap Mass Analyzer
Notes: Add in MS1 scans here

## Slide 8: Comprehensive MS/MS Acquisition
20 20 m/z-wide windows = 400 m/z
m/z
500
900
Time
~2 seconds (Q-Exactive)
MS Scan
Notes: Add in MS1 scans here

## Slide 9: Why????

## Slide 10: DIA Combines the Merits of DDA and Targeted Acquisition
at cost of precursor selectivity
Data Dependent Acquisition (DDA)
Parallel Reaction Monitoring (PRM)
Peptide Identification
Peptide Quantification
Data Independent Acquisition (DIA)
NLPFSVENK
NLPFSVENK
NLPFSVENK
NLPFSVENK
Database Searching
Targeted Analysis
Cycle time
Cycle time
Ying Sonia Ting

## Slide 11: Data Independent Acquisition
“All” Reaction Monitoring (coined by Bruce Lab)
20 20 m/z-wide windows = 400 m/z
m/z
500
900
Measure all fragment ions for all precursors between 500 and 900 m/z
Orbitrap
Linear Ion Trap
TOF
FT-ICR

## Slide 12: DIA Data Acquisition
m/z
500
900
Time
Comprehensive MS1 and MS2 Sampling
20 m/z
Un-Targeted
Un-Biased
Wide Isolation
500
m/z
520
m/z
Consistent Sampling
Chromatographic Time Scale
2
sec
Notes: Fix the Un-Targeted part, make it show the center of the scan, and then indicate how the peak is offset, and then have the peak jump around

## Slide 13: Targeted Chromatogram Extraction for VLENTEIGSDSIFDK++ (790.4 m/z)
m/z
500
900
780m/z
800
m/z
51.2
51.3
51.4
51.5
51.6
51.7
Retention Time
400
600
800
1000
1200
1400
m/z
0
1
2
3
4
Intensity x 10-5
Notes: Show query peptide here,

## Slide 14: DIA Data Analysis
Select Target Peptides
MS1
MS2
Extract Chromatograms
Select Chromatographic Peak
Set Peak Boundaries

## Slide 15: “SRM” on “All”
PRM/SRM
DIA

## Slide 16
DIA – A “complete” dataset
Query for any peptide/protein
SNPs
PTMS
Protein isoforms
Notes: Make the grey bars darker

## Slide 17: Is it Really SRM/PRM on All?
What are the drawbacks?
Notes: -- lack of charge state optimization
-- wide isolation windows causes myriad problems
-- you still have to choose a m/z range to target
-- cost
-- sensitivity
-- data acquisition downstream

## Slide 18
DIA
SRM
Isolation
Targeted
Narrow
Un-Targeted
Wide
Mass Analysis
Mass Filter
Low Resolution
Full Scan
HR/AM
20 m/z Isolation Window
0.7 m/z Isolation Window
Targeted Acquisition
Targeted Extraction

## Slide 19: The Advantages of DIA

## Slide 20
Data Acquisition
Peptide Selection
Chromatogram Extraction
Peak Integration
Transition Refinement
Hypothesis Generation
Hypothesis Test
Data Acquisition
Peptide Selection
Hypothesis Generation
Hypothesis Test
Transition Selection or Refinement
RT Scheduling
Peak Integration
DIA
SRM

## Slide 21
Data Acquisition
Peptide Selection
Chromatogram Extraction
Peak Integration
Transition Refinement
Hypothesis Generation
Hypothesis Test
Data Acquisition
Peptide Selection
Hypothesis Generation
Hypothesis Test
Transition Selection or Refinement
RT Scheduling
Peak Integration
DIA
SRM

## Slide 22
Data Acquisition
Peptide Selection
Chromatogram Extraction
Peak Integration
Transition Refinement
Hypothesis Generation
Hypothesis Test
Data Acquisition
Peptide Selection
Hypothesis Generation
Hypothesis Test
Transition Selection or Refinement
RT Scheduling
Peak Integration
DIA
SRM

## Slide 23
Data Acquisition
Peptide Selection
Chromatogram Extraction
Peak Integration
Transition Refinement
Hypothesis Generation
Hypothesis Test
Data Acquisition
Peptide Selection
Hypothesis Generation
Hypothesis Test
Transition Selection or Refinement
RT Scheduling
Peak Integration
DIA
SRM

## Slide 24: MS/MS Quantification is More Sensitive than MS
~30 bovine peptides spiked into a yeast background (Q-Exactive)

## Slide 25
150 peptides spiked in at 1 and 2 fmol (human neuronal cell background)

## Slide 26
150 peptides spiked in at 1 and 2 fmol (human neuronal cell background)

## Slide 27: Peptide Detection from DIA Data
How do you go about detecting peptides in DIA data?
  Can you use a database search?
  Different from SRM?
Are peptides being detected or identified?
Notes: Can a database searching approach work?

## Slide 28: Peak Picking is Challenging
Peptide: NLPFSVENK ++

## Slide 29: Detecting Peptides from DIA Data Directly
MS/MS spectra
Peptide spectrum matches
(PSMs)
p-value
q-value
p-value
q-value
Extracted MS/MS data
Evidence for detection
Spectrum-centric analysis
Peptide-centric analysis
APTNVTCILK
Peptides of interest
Protein sequence DB
What peptides best explain the data?
Which peptides are detected in our data?
example tools:
OpenSWATH, Skyline
example tools:
SEQUEST, MASCOT … etc.
Slide by Sonia Ting
p values are returned for every queried peptide
Notes: Instead of searching spectrum over a protein sequence database, as we usually do, we search targeted peptides over our data, trying to find whether this peptide was detected by our data (spectra collected).
Most database search algorithms assume one peptide per MS2 spectrum. This is how they are built.
Look at peptide identification problem in different light!
"We have a set of spectra and we would like to know what peptides that generated them".
"Which of these peptides do we have some kind of spectral evidence for?"

## Slide 30: Tools for Detecting Peptides in DIA Data are Becoming More Prevalent
 | No Spectral Library | Spectral Library
Peptide Centric | FT-ARM PECAN | OpenSWATH PeakView EncyclopeDIA Spectronaut
Spectrum Centric | DIA-UMPIRE Group-DIA MSE MSPLIT-DIA | 

## Slide 31: PECAN First Analyzes each Isolation Window Independently
Peptides of interest
PECAN
DIA
870-880 m/z
870-880 m/z
Analyze isolation window: 870-880 m/z
VTSAYLQDIENAYKK
Target peptides
Decoy peptides
KYANEIDQLYASTVK
Ying Sonia Ting

## Slide 32: PECAN Step 1: Peptide Scoring
Distribution of predicted fragment (b-/y-) ions
Fragment ion m/z
Frequency of occurrence
Protein sequence DB
PECAN
1a. Calculate fragment ions specificity
870-880 m/z
R[y1]: 175.1190
K[y1]: 147.1128
Peptides of interest
Ying Sonia Ting

## Slide 33: PECAN Step 1: Peptide Scoring
Protein sequence DB
PECAN
1b. Generate scoring vector for peptides
Peptides of interest
( for all b-/y- ions )
PECAN generated scoring vector
Target
Decoy
Ying Sonia Ting

## Slide 34: PECAN Step 1: Peptide Scoring
PECAN
1c. Score with fragment ion XICs
Peptides of interest
DIA
Centroid .mzML
time
transition
Target
Ying Sonia Ting

## Slide 35: PECAN Step 1: Peptide Scoring
PECAN
1c. Score with fragment ion XICs
Peptides of interest
DIA
Target
Decoy
Target
Decoy
Centroid .mzML

## Slide 36: PECAN Step 2: Background Subtraction
Protein sequence DB
PECAN
2a. Estimate background score
870-880 m/z
Thousands of randomly shuffled peptides
Scoring vectors
Ying Sonia Ting

## Slide 37: PECAN Step 2: Background Subtraction
Target
Decoy
PECAN
2b. Subtract background score
Peptides of interest
Target
Decoy
Target
Decoy
Ying Sonia Ting

## Slide 38: PECAN Step 3: Peak Picking
PECAN
3a. Pick the top N peaks (N=1)
Peptides of interest
1
2
3
4
5
6
Ying Sonia Ting

## Slide 39: PECAN Step 3: Peak Picking
PECAN
3b. Generate auxiliary scores for a peak
Peptides of interest
Fragment ions
- raw score
- background subtracted score
- spectrum similarity
- number of contributing ions
- mass error mean
- mass error variance
- rank
Precursor ions
- idotp (isotope dot product)
- mass error mean
- mass error variance
Peptide sequence
- charge state
- peptide length
- SSR hydrophobicity index
Ying Sonia Ting

## Slide 40: PECAN Step 4: Statistical analysis
PECAN
4. Statistical analysis
Peptides of interest
Decoy peptides
CVITNLAPTK
p-value
q-value
Percolator
Percolator, Käll et al. 2007
890-900m/z
400-410m/z
410-420m/z
880-890m/z
420-430m/z
APTNVTCILK
Peptides of interest

## Slide 41: Peptide Detection Sensitivity Scales with Isolation Width
Chart: MS/MS Isolation Width (m/z)
| 10 | 8 | 5 | 4 | 3 | 2
1 Injection | 9678 | 13104 | 14782 | 18567
2 Injections | 17847 | 20701 | 24017 | 25431
Pierce HeLa digest
Fusion Lumos
500 – 700 m/z
2 second cycle time
4 second cycle time
2 second cycle time
4 second
cycle time

## Slide 42: Isolation Window Width
Vs.
Vs.
2 m/z
10 m/z
20 m/z
DDA
DIA
Lower precursor selectivity
More peptides co-fragmented
More complex MS/MS spectra
More interference
Notes: Show the spectrum, and the chromatogram
Even with DDA >30% of the spectra are mixed
Quantitation can be limited by interference, show 10 m/z vs. 20 m/z (from overlap experiment)
Spectrum and chromatogram from loq data on a yeast background peptide

## Slide 43: Precursor Selectivity
2 m/z
ANFQGAITNR
Notes: Show the spectrum, and the chromatogram
Even with DDA >30% of the spectra are mixed
Quantitation can be limited by interference, show 10 m/z vs. 20 m/z (from overlap experiment)
Spectrum and chromatogram from loq data on a yeast background peptide

## Slide 44: Precursor Selectivity
10 m/z
ANFQGAITNR
Notes: Show the spectrum, and the chromatogram
Even with DDA >30% of the spectra are mixed
Quantitation can be limited by interference, show 10 m/z vs. 20 m/z (from overlap experiment)
Spectrum and chromatogram from loq data on a yeast background peptide

## Slide 45: Precursor Selectivity
20 m/z
ANFQGAITNR
Notes: Show the spectrum, and the chromatogram
Even with DDA >30% of the spectra are mixed
Quantitation can be limited by interference, show 10 m/z vs. 20 m/z (from overlap experiment)
Spectrum and chromatogram from loq data on a yeast background peptide

## Slide 46: Precursor Selectivity
Intensity
4e7
Retention Time (min)
25
26
10 m/z
ANFQGAITNR
Notes: RT 25-26 intensity 40e6

## Slide 47: Precursor Selectivity
Intensity
4e7
10 m/z
Retention Time (min)
Intensity
4e7
25
26
20 m/z
ANFQGAITNR
X

X
X


Notes: RT 25-26 intensity 40e6

## Slide 48: m/z Range Covered
Notes: EXTRA MATERIAL? – how this was calculated
by Jesse, also show numbers from Rapa study in overlap paper

## Slide 49: Cycle Time
20 20 m/z-wide windows = 400 m/z
m/z
500
900
~30 seconds
40 10 m/z-wide windows = 400 m/z
m/z
500
900
~30 seconds
10 Hz
15 scans
~7 scans
Cycle Time: 2 seconds
Cycle Time: 4 seconds
Notes: Add in MS1 scans here

## Slide 50: Cycle Time
20 20 m/z-wide windows = 400 m/z
m/z
500
900
~6 seconds
10 Hz
10 40 m/z-wide windows = 400 m/z
m/z
500
900
Notes: Add in MS1 scans here

## Slide 51: Number of Injections
~30 seconds
80 5 m/z-wide windows = 400 m/z
m/z
500
900
m/z
500
900
40 5 m/z-wide windows = 200 m/z
40 5 m/z-wide windows = 200 m/z
~30 seconds
Notes: What if I want to keep the better Cycle Time for quantification

## Slide 52: Improving Precursor Selectivity

X

X
Notes: Take a normal DIA scheme, with 20 m/z wide windows, overlap them (schematic figure)

## Slide 53: 20 m/z DIA
GVMNAVNNVNNVIAAAFVK

## Slide 54: 20 m/z DIA + Overlap
GVMNAVNNVNNVIAAAFVK

## Slide 55: 20 m/z DIA + Overlap + Demultiplexing
GVMNAVNNVNNVIAAAFVK

## Slide 56: 10 m/z DIA
GVMNAVNNVNNVIAAAFVK

## Slide 57: The Promise of DIA
Acquire a “molecular image” or “digital archive” of the sample
  Mine it over and over
  No scheduling
Direct queries (and p-values) for peptides of interest
Better quantitation than DDA
Notes: Acquire a molecular image
Go back to it whenever you want, requery
Better quantification when you do query

## Slide 58: Image Quality

## Slide 59: DDA vs. DIA vs. SRM
DDA
SRM/PRM
MS/MS Quantification
High Selectivity MS/MS
Quantification
MS Quantification
on almost everything
DIA
on targets
on semi-random subset
Notes: As instruments scan faster, DDA turns into DIA

## Slide 60: DIA for P100
Build a “super library” from a pooled sample
Map detections from library to all future single-shot DIA datasets

## Slide 61
Multiple
DDA Analyses
Pooled Sample
Database Search
Identification
Build a Spectral Library
Chromatogram Extraction
Quantification
1 DIA Run
20 m/z Isolation
/per sample
Acquire DIA Data
12 DIA Runs
2 m/z Isolation
Pooled Sample
Peptide
Detection
Build a Chromatogram Library
Spectral Library Workflow
Chromatogram
Notes: -- bad way to pick peptides
-- whole point of DIA is to be able to query for anything…

## Slide 62
DIA Library Generation with 12 LC-MS/MS Runs
12 LC-MS/MS Runs
400 m/z
1000 m/z
Inj. 12
Inj. 1
Inj. 2
25 x 2 m/z
MS/MS Scans
30k R.P.
400 m/z
450 m/z
MS1 30k R.P.
2.5 second
cycle time
Peptide Detection
…
…
…
Spectrum
Mill
Notes: Mention that 2 m/z isolation windows give great sensitivity for detection, getting close to PRM
Can extract chromatograms for any peptide, need a way of confirming its identity

## Slide 63: 12 LC-MS/MS Runs to Increase Sensitivity in All Future Runs
Chromatogram Library
12 LC-MS/MS Runs
2 m/z windows
Sample pool
1 LC-MS/MS Run
22 m/z overlapping windows
Single sample
10,571 phosphopeptides

## Slide 64: 12 LC-MS/MS Runs to Increase Sensitivity in All Future Runs
Chromatogram Library
12 LC-MS/MS Runs
2 m/z windows
Sample pool
EncyclopeDIA*
*Searle BC, et. al. In preparation
10,571 phosphopeptides
7,594 phosphopeptides

## Slide 65: Transition Refinement by EncyclopeDIA
Brian Searle
Raw Transition
Intensity
Retention Time (min)
Normalized Signal
Retention Time (min)
Retention Time (min)
Picked Transition
Intensity
quant threshold

## Slide 66: Peptide Detection in Multiple Runs
What’s happening here?
Notes: Detection in different places between runs

## Slide 67: Outlier Detection

## Slide 68: Outlier Detection: RT Alignment
1
2
Reference
4
3
5
30
…
Runs to Align
Reference RT
Run to Align RT

## Slide 69: Outlier Detection: Pairwise Deviation
1
3
2
4
30
…
2
3
4
30
…
# of Peptides
3σ
Fit to inner
98%
Retention Time Error (min)
3σ
3σ
3σ
3σ
3σ
…
…
…

## Slide 70: Outlier Detection
For each peptide…
LC-MS/MS Run
1
2
3
4
5
6
7
8
9
10
11
12
13
14
15
16
17
18
19
20
Aligned Retention Time

## Slide 71: Outlier Detection
For each peptide… sort by RT
LC-MS/MS Run
1
2
3
4
5
6
7
8
9
10
11
12
13
14
15
16
17
18
19
20
Aligned Retention Time
3σ ?

## Slide 72: Outlier Detection
Before
After
Retention Time
Retention Time
Retention Time

## Slide 73: Peak Boundary Imputation
100%: Detected
50%: Detected
1%: Imputed

## Slide 74: Imputation by Non-Parametric Pairwise Alignment
Retention Time (Run 1)
Retention Time (Run 2)
Retention Time (Run 1)

## Slide 75: P6000!!

## Slide 77: Acknowledgements
MacCoss Lab
Mike MacCoss
Sonia Ting
Brian Searle
Austin Keller
Richard Johnson
Lindsay Pino
Brendan MacLean
Han-Yin Yang
Gennifer Merrihew
Broad Institute
Jake Jaffe
Sebastian Vaca
Adam Officer
Caitlin Feeney
Amanda Creech
Jinal Patel
Xiaodong Lu
Dan Lam
Desiree Davidson

## Slide 78: Hackathon
Data: http://bit.ly/PCCSEData
How many of your probes from the previous hack-a-thon are found in the DIA data?
For each probe: what percentage of runs give non-zero values?